## Revision Tracking
| Current Version | Score Card_ GM_Lab_2017_R1.02
| The latest version of the score card is available for download, pls download online |  | Download here
History of amendments
S/N | Version No. | Remarks/Brief Description of changes | Effective date
1 | Score Card_ GM_Lab_2017_R1.01 | Uploaded first version | 2018-12-24
2 | Score Card_ GM_Lab_2017_R1.02 | Revision for user friendliness | 2020-06-15
BCA Green Mark for Laboratories
GM Lab: 2017
GM Lab 2017 Score Card

## Cover Page
GREEN MARK FOR LABORATORIES: 2017
[8 empty] | Please select | New | BSL1 |  | Date of Assessment: |  | Date of presentation:
|  | BCA GREEN MARK CERTIFICATION SCHEME | [5 empty] | Existing | Existing | BSL2 |  | Panel:
To | : | GD/ES |  |  |  |  | 2017-07-30 | Chemical |  | BSL3 |  | Category |  |  | Point Allocation | Point Scored
Thr’ | : | D/GMD(ND) | [5 empty] | CSDL2 | Biological | BSL4 |  | 1. Sustainable Design |  |  | 25 | 0
From | : | SM/GMD(ND) Leow Yock Keng (Lead Assessor) | [5 empty] | Gold | Chemical | CSDL0 |  | 1.1 | Leadership |  | 3 | 0
[9 empty] | Physical | CSDL1 |  | 1.1a | Organisation Governance |  | 3 | 0
Project Name | [9 empty] | CSDL2 |  | 1.2 | Design and Operation |  | 18 | 0
Reference No |  | GM2632/3/2017 | [6 empty] | Initial Assessment | CSDL3 |  | 1.2a | Sustainable Laboratory Design |  | 3 | 0
Address | [8 empty] | Recertification | CSDL4 |  | 1.2b | Sustainable Laobratory Operation |  | 6 | 0
Standards |  | Green Mark for Laboratories: 2017 | [7 empty] | NA |  | 1.2c | Environmental Credential of Project Team |  | 6 | 0
Type of Laboratories |  | Existing Chemical Laboratory | [6 empty] | Platinum |  |  | 1.2d | User Engagement |  | 3 | 0
|  | BSL2 | [6 empty] | GoldPlus |  |  | 1.3 | Urban Harmony |  | 4 | 0
Attached Report(s) |  | Green Mark report, Project Factsheet, Project Mock Certificate | [6 empty] | Gold |  |  | 1.3a | Sustainable Commuting |  | 2 | 0
Type of Audit |  | Initial Assessment | [5 empty] | Section |  | Point scored | Max | 1.3b | Greenery Provision |  | 2 | 0
[8 empty] | Advanced Green Effort |  | 0 | 5 | 2. Laboratory Energy Performance |  |  | 35 (+5) | 1
[6 empty] | Basic Lab information |  | Smart and Healthy Laboratory |  | 0 | 20 | 2.1 | Energy Efficiency |  | 20 (+4) | 0
[6 empty] | No of modular labs | 0 | Resource Stewardship |  | 0 | 20 | 2.1a | Air Conditioning System Efficiency |  | 8 | 0
[6 empty] | Floor area | 0 | Laboratory Energy Performance |  | 1 | 20 |  | ~ 2 additional credits for Physical Laboratories |  | (+ 2)
[8 empty] | Sustainable Design |  | 0 | 20 | 2.1b | Ventilation Optimisation |  | 6 | 0
[12 empty] | 2.1c | Lighting Efficiency |  | 6 | 0
[13 empty] | ~ 1 additional credits for Physical Laboratories |  | (+ 2)
[12 empty] | 2.2 | Energy Effectiveness |  | 15 (+1) | 1
[12 empty] | 2.2a | Factilites Energy |  | 1 | 0
The assessment for the abovementioned project has been conducted and the project is: | [11 empty] | 2.2b | Fume Hood/Biosafety Cabinet Performance |  | 4 | 0
þ | Recommended for Gold Award | [10 empty] | 2.2c | Plug Load Management |  | 4 | 0
¨ | Not Recommended for Award | [11 empty] | ~ 1 additional credits for Physical Laboratories |  | (+1)
[12 empty] | 2.2d | Equipment Maintenance |  | 2 | 0
Reason(s) : | [11 empty] | 2.2e | Energy Efficient Practices, Design and Features |  | 4 | 1
1. | The project scores more than 50 points, eligible for Gold award. There was an delay of certification of this project from the previous 3 pilot projects as the team want to enhance internal SOPs to achieve GM Gold rating. | [10 empty] | 3. Resource Stewardship |  |  | 25 | 0
[12 empty] | 3.1 | Water |  | 8 | 0
[12 empty] | 3.1a | Water Efficient Measures |  | 5 | 0
[12 empty] | 3.1b | Water Usage Monitoring |  | 1 | 0
[12 empty] | 3.1c | Water Efficiency Improvement Plans |  | 1 | 0
2. | The project met minimum requirement to achieve Green Mark Gold rating | [10 empty] | 3.1d | Alternative Water Soucres |  | 1 | 0
[12 empty] | 3.2 | Materials |  | 11 | 0
Key issues / concerns |  |  | Panel's Comments |  |  | Additional information required | [5 empty] | 3.1a | Sustainable Construction |  | 2 | 0
None |  |  | NA | [8 empty] | 3.1b | Sustainable Products |  | 7 | 0
[12 empty] | 3.1c | Material Management |  | 2 | 0
[12 empty] | 3.3 | Waste |  | 6 | 0
[12 empty] | 3.1a | Operatiopnal Waste Management |  | 6 | 0
Submitted and Checked by: |  |  |  |  | Checked by: | [6 empty] | 4. Smart and Healthy Laboratory |  |  | 25 | 0
[12 empty] | 4.1 | Indoor Air Quality |  | 15 | 0
[12 empty] | 4.1a | IAQ Performance |  | 8 | 0
[12 empty] | 4.1b | Contaminants Management |  | 7 | 0
Leow Yock Keng (Lead Assessor) |  |  |  |  | Yong Siew Hwa (Co-Assessor) | [6 empty] | 4.2 | Spatial Quality |  | 12 | 0
SM, Green Mark Department (New Development) |  |  |  |  | PM, Green Mark Department (Existing Buildings) | [6 empty] | 4.2a | Lighting |  | 6 | 0
Signature / Date |  |  |  |  | Signature / Date | [6 empty] | 4.2b | Acoutics |  | 1 | 0
[12 empty] | 4.2c | Wellbeing and Safety |  | 5 | 0
[12 empty] | 4.3 | Smart Building Operations |  | 8 | 0
[12 empty] | 4.3a | Energy Monitoring |  | 2 | 0
[12 empty] | 4.3b | Demand Control |  | 4 | 0
Director, GMD(ND) |  |  |  | Group Director, ESG | [7 empty] | 4.3c | Integration and Analytics |  | 2 | 0
Signature/Date |  |  |  | Signature/Date | [7 empty] | 5. Advanced Green Efforts |  |  | 20 | 0
[12 empty] | 5.1 | Implementation of good practices |  | 10 | 0
[12 empty] | 5.2 | Innovative Green Features |  | 10 | 0
[12 empty] | All design must not compromise safety |  | Total | 140 | 1

## Scoring
Item |  | Assessment Evidence/ Comments |  | Points | Status | Field | Fill-in | [11 empty] | Item |  | Assessment Evidence/ Comments |  | Points | Status | Field
1. SUSTAINABLE DESIGN | [7 empty] | KIV |  | Choice | [8 empty] | 4. IAQ PERFORMANCE
1.1 | Leadership |  |  |  | Assessor's use |  |  | Inadequate |  |  | i | ii | iii | iv | v |  |  |  | 4.1 | Indoor Air Quality |  |  |  | Assessor's use
1.1a | Organisation Governance | · | [7 empty] | Not scoring | Establish key appointment holders for laboratories operation with organisation chart made known to lab users | Regular update to senior management | Polices review involving senior management | Senior management committed to environmental policy |  |  |  |  | 4.1a | IAQ Performance | · |  |  |  |  | Not scoring | Conduct IAQ audit once every three years (Indicative method) | Conduct IAQ audit once every three years (Reference method) | Conduct yearly post occupancy evaluation | Implement corrective actions | Optimising airflow based on CFD or physical modeling | Develop an active IAQ management programme
[27 empty] | 1 | 2 | 2 | 1 | 2 | 1
1.2 | Design and Operation
1.2a | Sustainable Laboratory Design | · | [7 empty] | Not scoring | Integrative design process | Performs design/audit and ROI evaluation prior to renovation | Conduct regular laboratory infrastructure, energy and hazard assessment | Permanent committee to evaluate laboratory for continual improvement |  |  |  |  | 4.1b | Contaminants Management | · |  |  |  |  | Not scoring | Use of MERV 13 filter | Use of MERV 14 filter | Use of HEPA filter | Contaminants monitoring (Spot measurement) | Contaminants monitoring (Continuous measurement) | Fume Hood Commissioning (Tier 1) | Fume Hood Commissioning (Tier 2) | Biosafety Cabinet Comissioning (Tier 1) | Biosafety Cabinet Comissioning (Tier 2)
[27 empty] | 1 | 2 | 1 | 1 | 2 | 1 | 2 | 1 | 2
1.2b | Sustainable Laboratory Operation | · | Not scoring | 0 |  | Name |  |  |  | Not scoring | Laboratory is located in a Green Mark Platinum Building | Laboratory is located in a Green Mark GoldPLUS Building | Laboratory is located in a Green Mark Gold Building | Laboratory demontrates % of energy saving over laboratory's total energy consumption
[6 empty] | % of energy saving |  |  |  |  | 6 | 4 | 3 | 0 |  |  |  |  | 4.2 | Spatial Quality
[19 empty] | 4-2a | Lighting | · |  |  |  |  | Not scoring | Provision of daylighting | Provision of daylighting with glare control strategies | Lighting Zoning with adequate switches | Adequate lighting level | Controllability of lighting system | High frequency ballasts for fluorescent fittings | LED with ≤ 30% flicker | High frequency ballasts for fluorescent fittings and LED with ≤ 30% flicker
1.2c | Environmental Credential of Project Team | · |  |  |  | Custom | [20 empty] | 1 | 2 | 1 | 1 | 1 | 1 | 1 | 1
(Please fill in the text and score required on the left)
1.2d | User Engagement | · | [7 empty] | Not scoring | Green guide for laboratory occupants | Lab users are informed about the energy usage | Guidelines or commuications to ensure chemicals/hazardous materials are used efficiently to minimise waste generation
[11 empty] | 1 | 1 | 1 | [5 empty] | 4.2b | Acoustics | · |  |  |  |  | Not scoring | Ambient sound level of ≤ 55 dBA without the lab equipment in operation. | Ambient sound level of ≤ 65 dBA with the lab equipment in operation.
[19 empty] | 4.2c | Wellbeing and Safety | · |  |  |  |  | Not scoring | Controllability of temperature | Provision of ergonomic (adjustable) chairs for users | Provision of ergonomic (adjustable) chairs and workstations for users | Display panel showing the level of CO2, TVOC and/or chemicals for each laboratory | Designated breakout areas for occupants to relax/eat outside laboratory environment | Space pressure control on monitoring system to prevent exfiltration of contaminants | Availability to relevant safety devices | Safety measures for performing lab tasks | Periodic briefing on lab safety and emergency procedures
1.3 | Urban Harmony
1.3a | Sustainable Commuting | · | [7 empty] | Not scoring | Provision of bus services to pool saff to work and/or after work | Provision of regular shuttle buses to ferry staff between premises during office hours | Encourage staff to use public transport, bicycle or other human-powered/electric personal transportation through awareness and incentive programs | Allocation of spaces to encourage sustainable commuting such as bicycle | Foster effective carpooling program (involving at least 15% of staff)
#VALUE!
1.3b | Greenery Provision | · | [7 empty] | Not scoring | Provision of access to greenery at common spaces | Provision of planters or plotted plants, equivalent to 0.1% of lab areas. | Provision of planters or plotted plants, equivalent to 0.2% of lab areas. | [5 empty] | 4.3 | Smart Building Operations
[11 empty] | 0.5 | 1 | 2 | [5 empty] | 4.3a | Energy Monitoring | · |  |  |  |  | Not scoring | Provision of monitoring of individual electrical circuits/equipment | Provision of monitoring of electrical loads with link to Building Management System (BMS) or equivalent
2. LAB ENERGY PERFORMANCE | [5 empty] | Mode of measurements | Measured readings | 1
2.1 | Energy Efficiency |  |  |  |  | Chiller System Efficiency in kW/ton |  | 100 | 9 |  | DCS system | Physical Lab | Natural Ventilation | Physical Lab with DCS system |  |  |  |  | 4.3b | Demand Control | · |  |  |  |  | Not scoring | Provision of automated night-setback for lower ACH for laboratory during non-operation hours | Automatic switch off/dim lighting during non-occupied hours | Provision of variable airflow control for different zones | Admin areas controlled by separate demand control ventilation
2.1a | Air Conditioning System Efficiency | · | Not scoring | 0 |  | Mode of measurements | Measured readings | 1 |  | kW/ton | m/s
[6 empty] | Air distribution metric | kW/ton | 100 |  | Measured readings | Nameplates
[6 empty] | Air distribution | 0
2.1b | Ventilation optimisation | · |  |  |  | Pls select (if applicable) |  | Not inputted yet |  | Not scoring | Adopt risk based assessment to minimise lab ventilation rate | Segregation of areas with different ACH design for more than 70% of applicable areas | Localised exhaust or provision of effective isolation zoning of heat-generating equipment for 70% of critical heat sources and instrument | Reuse of return or recirculated cooled air from office/support areas for air make-up to laboratories areas or design for air recirculation within laboratory while meeting safety requirement. | Apply run-around coil, wraparound heat pipe or sensible heat recovery between the exhaust and fresh airstreams. Heat recovery from exhaust systems to pre-cool outdoor air | Use of onsite recovered energy to reheat air | Adopt VAV lab air flow and variable flow exhaust controls (Variable design) | Multi-stack exhaust plenum with staged-exhaust fans (Shut-off design)
[19 empty] | 4.3c | Integration and Analytics | · |  |  |  |  | Not scoring | Provision of website and/or accessible monthly readout per laboratory | Provision of energy portal and/or dashboard for laboratory management team
5. ADVANCED GREEN EFFORTS
[19 empty] | 5-1a | Implementation of good practices | · | Not scoring | 0 |  | Custom
[19 empty] | 5-1b | Innovative Green Features | · | Not scoring | 0 |  | Custom (Please fill in the text and score required on the left)
2.1c | Lighting Efficiency | · | Not scoring | 0 |  | % improvement over baseline
2-2a | Facilites Energy | · |  | 0 | [5 empty] | Not scoring | Setting target improve laboratories energy performance
2-2b | Fume Hood Performance | · | [7 empty] | Not scoring | VAV fume hood | VAV fume hood with occupancy sensor or sash position monitoring | VAV fume hood with face velocity at 0.54m/s | Sash position monitoring linked to BMS
[6 empty] | Face Velocity | 0.54 |  |  |  | 0.5 | 1 | 0.6 | 1
[10 empty] | Not scoring | Proper heat load calculation with diversity and occupancy consideration or benchmarking of receptacle load | Provision of 2 energy efficient labelled equipment | Use of Tier 1 energy efficient freezers for more than 80% of freezers | Active plug load management based on operation schedule | Overnight equipment management system | Freezer does not meet requirement. Ultra low freezer at 14.7kWh/litre/year
2.2c | Plug load management and operation | · |  |  |  | No of type of energy efficient equipment | 2 |  |  |  | 1 | 1 | 1 | 1 | 1 |  | 0
[6 empty] | Freezer | Tier 1 |  |  | Freezer does not meet requirement. Ultra low freezer at 14.7kWh/litre/year
[10 empty] | 1 | 2 | 3 | 4
[10 empty] | Tier 1 | Tier 2
2.2d | Equipment Maintenance | · | [7 empty] | Not scoring | Effective regular maintenance of equipment | Inventory optimisation
2.2e | Energy Efficient Practices, Design and Features | · | EUI = XX kWh/m²/yr, EEI = XX kWh/m²/yr | 1 |  | Custom
[6 empty] | (Please fill in the text and score required on the left) |  |  |  | Demand control ventilation based on active air quality sampling | Ductless fume hoods for at least 80% of applicable equipment | Ductless biosafety cabinets with HEPA filter for at least 80% of applicable equipment | Adoption of task lightings to reduce general lighting W/m2 design (achieving overall W/m2 with at least 40% improvement overcode requirement)
[10 empty] | 2 | 1 | 0.5 | 1
3. RESOURCE STEWARDSHIP
3.1 | Water
3-1a | Water Efficient Measures | · | [7 empty] | Not scoring | Provision of products certified under WELS | Achieve PUB Water Efficient Building Certification (Basic) | Achieve PUB Water Efficient Building Certification (Silver/Gold) | Provision of water efficient laboratory equipment | Recyce water generated from regular flushing of laboratory safety fitting
[12 empty] | 2 | 3 | 1 | 1
[10 empty] | Not scoring | Use of local sub-meter for equipment | Use of local sub-meter with leak detector for equipment | Use of smart meters with remote monitoring for equipment
3.1b | Water Usage Monitoring | · | [8 empty] | 0.5 | 0.5 | 1
3.1c | Water Efficiency Improvement Plans | · | [7 empty] | Not scoring | Implement water efficiency improvement plans with target setting
3.1d | Alternative Water Sources | · |  |  |  | On-site recycled water | AHU condensate |  |  | Not scoring | Use of NEWater | Use of rainwater harvested | Use of AHU condensate
3.2 | Materials
3.2a | Sustainable Construction | · | [7 empty] | Not scoring | Conserve at least 50% of existing finishing in renovation
3.2b | Sustainable Products | · |  |  |  | Custom
(Please fill in the text and score required on the left)
3.2c | Material management | · | [7 empty] | Not scoring | Labelling with details of contents, ownership and (where relevant) hazard and emergency details | Records of equipment for tracking and ensure surplus stock can be re-distributed and to reduce unnecessary procurement.
3.3 | Waste
3.3a | Operational waste management | · | [7 empty] | Not scoring | Provision of recycling facilities at a central location | Provision of recycling facilities at every laboratory/stragetic locations | Recycling of equipment packaging waste | Provide laundry service for lab coat to reduce wastage and encourage hygiene and safety. | Waste collecting and monitoring including hazard waste management | Waste management improvement plan
[11 empty] | 1 | 2 | 1 | 1 | 1 | 1

## Lab EEI
Energy Efficiency Index (EEI)
Electricity Consumption (recent 12 months)
Month | Total Consumption / kWh
Jun 18
Jul 18
Aug 18
Sep 18
Oct 18
Nov 18
Dec 18
Jan 19
Feb 19
Mar 19
Apr 19
May 19
Total | 0
Average | 0
GFA Breakdown for Laboratory | 0
GFA Breakdown for Office, Facilities Support Room | 0
GFA Breakdown for Lan Room 24 hours operation | 0
Operating Hours (hrs/wk) | 168
Total Energy Consumption (recent 12 months) | 0
Energy Consumption for Office and Lan room, support areas (recent 12 months) | 0 | *Note - Using average of 188kWh/m2/yr for office energy consumption, extracted from BCA source. Approximate that exclusion of water-side is 65% of total consumption. Lan room is taken as 270W/m2. Majority of lan room consumption is sensible load from equipment. | A/C is usu 50% of energy consumption and water-side is 70% of A/C consumption. Therefore, consumption excluding waterside is 65%. Assume lan room's heat are 80% from equipment sensible load. With A/C COP of 6, hence, 80% x 270 + 20% x 270/6 x 30% = 218.7 (which is 80% of original load)
Energy Consumption for Laboratory (recent 12 months) | 0
Laboratory EUI |  | kWh/m2/yr
Laboratory EEI |  | kWh/m2/yr

## Additional Info
Physical Properties (Pls enter value without unit) |  |  | Operation Properties (Pls enter value without unit)
Type of Laboratory and safety level |  |  | Space Temp setpoint (deg C) | [7 empty] | FALSE
|  |  | Space RH setpoint (%) | [7 empty] | FALSE
|  |  | Lab annual power consumption (kWh) | 0
Total Floor area (m2) A+B |  |  | Operating Hours/Week (hr/wk) | 168
Lab Floor area (m2) A |  |  | EUI (kWh/m2/yr)
Supporting/ Common area (m2) B |  |  | EEI (normalised) kWh/m2/yr
Lab volume (m3) |  |  | M&E Performance (Pls enter value without unit)
Number of modular labs |  |  | ACH (Operating)
Number of fume hoods |  |  | ACH (Non-Occupancy)
Number of biosafety cabinets |  |  | Air distribution performance (kW/ton)
Number of freezers |  |  | Lighting LPD W/m2
Type of luminaires (majority)
|  |  | Lighting fitting details |  | Eg. 2x28W
BSL | CSDL | Safety Level | Total Floor area (m2) | Lab floor area | Supporting/Common area (m2) | Lab Volume (m3) | No. of modular labs | No. of fume hoods | No. of biosafety cabinets | No. of freezers | Space Temp setpoint | Space RH setpoint | Lab annual power consumption (kWh) | OH/Week | EUI (kWh/m2.yr) | EEI (normalised) (kWh/m2.yr) | ACH (Operating) | ACH (Non-Occupancy) | Air distribution performance (kW/ton) | LPD W/m2 | Type of luminaires | Lighting details
0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 168 |  |  | 0 | 0 | 0 | 0 | 0 | 0
BSL
CDSL
NA